- Annex 4. Item 5.1. – Chapter 1.1.5. Quality management in veterinary testing laboratories

- MEETING OF THE WOAH BIOLOGICAL STANDARDS COMMISSION

- Paris, 4–8 September 2023

4

5

6	C H A P T E R 1. 1.5 .

# 7	QU AL IT Y M ANA G EM ENT IN V ET ER INAR Y

8	T EST ING LAB OR AT OR IES

9	SUMMARY

- Valid laboratory results are essential for diagnosis, surveillance, and trade. Such results are achieved by the

- use assured through implementation of good a management practices, valid system that supports accurate

- and consistent test and calibration methods, proper techniques, quality control and quality assurance, all

- working together within a quality management system. Laboratory quality management includes technical,

- managerial, and operational elements of testing performing, interpreting and the interpretation of reporting

- test results. A quality management system enables the laboratory to demonstrate both competency and an

- ability to generate consistent technically valid results that meet the needs of its customers. The need for

- Mutual recognition and acceptance of test results for international trade, and the acceptance accreditation

- of tests to international standards such as ISO/IEC 25 17025:2005 (General Requirements for the

- Competence of Testing and Calibration Laboratories) (ISO/IEC, 2005 2017b) requires good suitable

- laboratory quality management systems. This chapter is not intended to reiterate the requirements of

- ISO/IEC 17025, nor has it been endorsed by accreditation bodies. Rather, it outlines the important issues

- and considerations a laboratory should address in the design and maintenance of its quality management

- system, whether or not it has been formally accredited regardless of formal accreditation status. Chapter

- 1.1.1 Management of veterinary diagnostic laboratories gives an introduction to veterinary diagnostic

- laboratories introduces the components of governance and management of veterinary laboratories that are

- necessary for the effective delivery of diagnostic services, and highlights the critical elements that should be

- established as minimum requirements.

## 28	A. KEY CONSIDERATIONS FOR THE DESIGN AND MAINTENANCE OF A

29	LABORATORY QUALITY MANAGEMENT SYSTEM

- To ensure that the quality management system is appropriate and effective, the design must be carefully thought out

- planned and, where accreditation is sought, must address all criteria of the appropriate quality standard. The major

- categories of considerations and the their associated key issues and activities within each of these categories are outlined

- in the following eight sections of this chapter.

- ISO/IEC: International Organization for Standardization/International Electrochemical Commission.

#### 1.	The work, responsibilities, and goals of the laboratory

- Many factors affect the necessary elements and requirements of a quality management system. These factors include,

- including:

- i)	Type of testing done performed, e.g. research versus diagnostic work;

- ii)	Purpose and requirements of the test results, e.g. for import or /export quarantine testing, surveillance, emergency

- disease exclusion, declaration of freedom from disease post-outbreak;

- iii)	Potential impact of a questionable or, erroneous or unfavourable result, e.g. detection of foot and mouth disease

- (FMD) in an FMD-free country;

- iv)	The tolerance level of Risk and liability tolerance, e.g. vaccination vs versus culling or /slaughter;

- v)	Customer needs (requirements, e.g. sensitivity and specificity of the test method, cost, turnaround time, strain or

- genotype level of characterisation), e.g. for surveillance, or declaration of freedom after outbreak;

- vi)	The role of the laboratory Role in legal work or in regulatory programmes, e.g. for disease eradication and declaration

- of disease freedom to the WOAH;

- vii)  The role of the laboratory Role in assisting with, confirming, or overseeing the work of other laboratories (e.g. as a

- reference laboratory);

- viii) Business goals of the laboratory, including the need for any third-party recognition or accreditation.

#### 2.	Standards, guides, and references

- The laboratory should choose reputable and accepted follow globally recognised standards and guides to assist in

- designing the quality management system. For laboratories seeking accreditation formal recognition of testing competency,

- and for all WOAH Reference Laboratories, the use of ISO/IEC 17025 (ISO/IEC, 2005 2017b) or equivalent will be is

- essential. This standard includes specifies managerial and technical requirements and accredited laboratories that are

- compliant are regarded as competent. Further information on standards may be obtained from the national standards body

- of each country, from the International Laboratory Accreditation Cooperation (ILAC) 26, and from accreditation bodies, e.g.

- the National Association of Testing Authorities (NATA), Australia, the United Kingdom Accreditation Service (UKAS), the

- American Association for Laboratory Accreditation (A2LA), etc. Technical and international organisations such as AOAC

- International (The Scientific Association Dedicated to Analytical Excellence; formerly the Association of Official Analytical

- Chemists) and the International Organization for Standardization (ISO) publish useful references, guides, application

- documents and standards that supplement the general requirements of ISO/IEC 17025. Other relevant documents may

- include guides and application documents providing interpretative criteria and recommendations for the application of

- ISO/IEC	17025	in	the	field	of	veterinary	testing	for	both	applicant	and	accredited	facilities,	e.g.

- https://www.nata.com.au/phocadownload/spec-criteria-guidance/animal-health/Animal-Health-ISO-IEC-17025-

- Appendix.pdf	https://nata.com.au/files/2021/05/Animal-Health-ISO-IEC-17025-Appendix-effective-March2021.pdf;

- Newberry & Colling, 2021.

- The ISO International Standard 9001 (ISO, 2015), is a certification standard specifies the requirements for quality

- management systems and while it may be a useful supplement framework to a underpin a laboratory quality system,

- fulfilment of its requirements does not necessarily ensure or imply assure technical competence (in the areas listed in

- Section 3 Accreditation). Conformance to the requirements of ISO 9001 is assessed by a certification body that is

- accredited to undertake such assessments by the national accreditation body to undertake such assessments. When a

- laboratory meets the requirements of ISO 9001, the term registration or certification is used to indicate conformity, not

- accreditation.

- With the advent of stronger alliances between medical and veterinary diagnostic testing under initiatives such as “One

- Health”, some laboratories may wish to choose to follow other ISO standards such as ISO 15189 Medical Laboratories –

- Requirements for Quality and Competence (ISO/IEC, 2012), which include 2022), for testing of human samples, e.g. for

- zoonotic diseases. It should be noted that for veterinary laboratories, limited availability of suitable material may render

- validation difficult; under these circumstances it is necessary to highlight the limited validation status when reporting results

- and their interpretation (Stevenson et al., 2021).

#### 3.	Accreditation

- If the laboratory decides to proceed with formal recognition of its a laboratory’s quality management system and testing,

- then is sought, third party verification of its conformity with the selected standard(s) will be is necessary. ILAC has published

- specific requirements and guides for laboratories and accreditation bodies. Under the ILAC system, ISO/IEC 17025 is to

- ILAC: The ILAC Secretariat, PO Box 7507, Silverwater, NSW 2128, Australia; http://ilac.org/

- be used for laboratory accreditation of testing or calibration activities. Definitions regarding laboratory accreditation may

- be found in ISO/IEC International Standard 17000: Conformity Assessment – Vocabulary and General Principles (ISO/IEC,

- 2004a 2020). Accreditation is tied to dependent on demonstrated competence, which is encompasses significantly more

- than having and following documented procedures. Providing a competent and customer-oriented service also means that

- the laboratory requires:

- i)	Adequate facilities and environmental controls;

- ii)	Has Appropriately qualified and trained personnel with a depth of technical knowledge commensurate with

- appropriate level of authority;

- iii)	Has appropriate Equipment with planned that is appropriately verified and managed in accordance with the relevant

- maintenance and calibration schedule;

- iv)	Has adequate facilities and environmental control;

- v)	Has procedures and specifications that ensure accurate and reliable results;

- vi)	Implements continual improvements in testing and quality management;

- vii)  Can assess the need for and implement appropriate corrective or preventive actions, e.g. customer satisfaction;

- viii) Accurately assesses and controls uncertainty in testing;

- ix)	Appropriate sample and materials management processes;

- x)	Has Technically valid and validated test methods, procedures and specifications that are, documented in accordance

- with the requirements of the applicable standard or guidelines, e.g. Chapter 1.1.6 Principles and methods of validation

- of diagnostic assays for infectious diseases and, chapters 2.2.1 to 2.2.8 Recommendations for validation of diagnostic

- tests and Special Issue of the Scientific and Technical Review (2021)27;

- xi)	Demonstrates Demonstrable proficiency in the applicable test methods used (e.g. by regular participation in

- proficiency tests on a regular basis testing schemes);

- xii)  Accurate assessment and control of the measurement of uncertainty in testing;

- xiii) Good documentation practices, e.g. ALCOA+ principles (i.e. Attributable, Legible, Contemporaneous, Original,

- Accurate, Complete, Consistent, Enduring, Available);

- xiv) Non-conformance management process, including detection, reporting, risk-assessment and implementation of

- effective corrective and preventive actions;

- xv)  Complaints management;

- xvi) Adequate control of data and information;

- xvii) Appropriate reporting and approval process;

- xviii) Culture of continual improvement.

- xix) Has demonstrable competence to generate technically valid results.

#### 4.	Selection of an accreditation body

- To facilitate the acceptance of the laboratory’s test results for trade, the accreditation standard used must be recognised

- by the international community and the accreditation body recognised as competent to accredit laboratories. Programmes

- for the recognition of accreditation bodies are, in the ILAC scheme, based on the requirements of ISO/IEC International

- Standard 17011: Conformity Assessment – General Requirements for Accreditation Bodies Accrediting Conformity

- Assessment Bodies (ISO/IEC, 2004b 2017a). Information on recognised accreditation bodies may be obtained from the

- organisations that recognise them, such as the Asia-Pacific Accreditation Cooperation (APAC), the Inter-American

- Accreditation Cooperation (IAAC), and the European Co-operation for Accreditation (EA).

- Accreditation bodies may also be signatory to the ILAC and regional (e.g. APAC) mutual recognition arrangements (MRAs).

- These MRAs are designed to reduce technical barriers to trade and further facilitate the acceptance of a laboratory’s test

- results in foreign markets. Further information on the ILAC MRA may be obtained from the www.ilac.org.

#### 5.	Determination of the scope of the quality management system or of the laboratory’s

- accreditation

- Available at: https://doc.woah.org/dyn/portal/index.xhtml?page=alo&aloId=41245

- The scope of the quality management system should cover all areas of activity affecting all include all activities that impact

- testing that is done at performed by the laboratory. Whilst only accredited laboratories are obliged to meet the requirements

- of the relevant standard as detailed below, these, the guiding principles should be considered best practise and are relevant

- to all testing laboratories.

- Laboratories accredited A laboratory’s accreditation to ISO/IEC 17025 have includes a specific list of those accredited tests

- that are accredited, called, referred to as the schedule or scope of accreditation or the scope. Veterinary testing facilities

- include government and private facilities, veterinary practices, university veterinary schools, and other laboratories for the

- testing of animals and animal products for the diagnosis, monitoring and treatment of disease. In principle, if new testing

- methods are introduced these must be assessed and accredited before they can be added to the scope, however a flexible

- scope can be implemented that assesses the laboratory as competent to add tests to scope, which are then formally added

- at the next accreditation visit. The quality management system should ideally cover all areas of activity affecting all testing

- that is done at the laboratory. However, it is up to the laboratory to decide which tests are to be accredited and included in

- the scope. If an accredited laboratory also offers unaccredited non-accredited tests, these must be clearly indicated as

- such on any reports that claim or make reference to accreditation. Factors It is ultimately the decision of the laboratory to

- decide which tests require inclusion in the scope of accreditation, and factors that might affect the laboratory’s choice of

- tests for scope of accreditation this decision include:

- i)	The impact of initial accreditation on resources within a given deadline;

- ii)	Associated risks and opportunities;

- iii)	Initial investment required (e.g. time, resources);

- iv)	A Contractual requirement for accredited testing (e.g. for international trade, research projects);

- v)	The Importance of the test and the potential impact of an incorrect result;

- vi)	The cost of maintaining an accredited test versus frequency of use;

- vii)  Availability of personnel, facilities and equipment;

- viii) Availability of appropriate materials and reference standards (e.g. standardised reagents, internal quality control

- samples controls, reference cultures) and

- ix)	Access to proficiency testing schemes;

- x)	The quality assurance control processes necessary for materials, reagents and media;

- xi)	The validation status, e.g. access to field samples from infected and non-infected animals, technical complexity and

- reliability of the test method;

- xii)  The Potential for subcontracting of accredited tests.

#### 6.	Quality assurance, quality control and proficiency testing

- Quality assurance (QA) is the part element of quality management focused on providing confidence that quality defined

- requirements will be are fulfilled. The requirements may be internal or defined in an accreditation or certification standard.

- QA is process-oriented and ensures provides the right things are being done in the right way appropriate inputs to prevent

- problems arising.

- Quality control (QC) is the systematic and planned monitoring of outputs to ensure the minimum levels of quality

- requirements have been met. For a testing laboratory, this is to ensure test processes ensures tests are working correctly

- performing consistently and reliably, and results are within the expected acceptable parameters and limits. QC is test

- orientated and ensures the results are as expected-oriented and ensures detection of any problems that arise.

- Proficiency testing (PT), sometimes referred to as external quality assurance or (EQA), is the determination assessment

- of a laboratory’s performance by when testing a standardised panel of specimens of undisclosed content. Ideally, PT

- schemes should be run managed by an external independent provider. Participation in proficiency testing schemes enables

- the laboratory to assess and demonstrate the their testing reliability of results by in comparison with those from other

- participating laboratories.

- All laboratories should, where possible, participate in external proficiency testing schemes appropriate to their testing.

- Participation the suite of tests provided; participation in such schemes is a requirement for accredited laboratories. This

- provides an independent assessment of the testing methods used and as well as the level of staff competence. If such

- schemes are not available, valid alternatives may be used, such as ring trials organised by reference laboratories, inter-

- laboratory testing, use of certified reference materials or internal quality control samples, replicate testing using the same

- or different methods, retesting of retained items, and or correlation of results for different characteristics of a specimen.

- Providers and operators of proficiency testing programmes should be accredited to ISO/IEC 17043 – Conformity

- Assessment – General Requirements for Proficiency Testing (ISO/IEC, 2010).

- Proficiency testing material from accredited providers has been is well characterised and any spare material, once the

- proficiency testing has been completed, can be useful to demonstrate staff competence or for test validation. Information

- about selection and use of reference samples and panels is available in Chapter 2.2.6 Selection and use of reference

- samples and panels. Proficiency testing and reproducibility scenarios are described by Johnson & Cabuang (2021) and

- Waugh & Clark (2021), respectively.

#### 7.	Test methods

- ISO/IEC 17025 requires the use of appropriate test methods and has requirements for their selection, development, and

- validation to show demonstrate fitness for purpose.

- This Terrestrial Manual provides recommendations on the selection of test methods for trade, diagnostic and surveillance

- purposes in the chapters on specific diseases. Disease-specific chapters include, or will include in the near future, a table

- of the tests available for the disease, graded against the test’s fitness for purpose; these purposes are defined in the WOAH

- Validation Template (chapter 1.1.6), which identifies six main purposes for which diagnostic tests may be carried out. The

- table is intended to be as a general guide to test application. ; the fact that a test is recommended does not necessarily

- mean that a laboratory is competent to perform it. The laboratory quality system should incorporate provision of evidence

- of competency.

- In the veterinary  profession laboratories, other standard methods (published in international, regional, or national

- standards) or fully validated methods (having undergone a full collaborative study and that are published or issued by an

- authoritative technical body such as the AOAC International) may be preferable to use, but may not be available. Many

- veterinary laboratories develop or modify methods, and most laboratories have test systems that use non-standard

- methods, or a combination of standard and non-standard methods. In veterinary laboratories, even with the use of standard

- methods, some in-house evaluation, optimisation, or validation is generally must be done required to ensure valid results.

- Customers and laboratory staff must have a clear understanding of the performance characteristics of the test, and

- customers should be informed if the method is non-standard. Many veterinary testing laboratories will therefore need to

- demonstrate competence in the development, adaptation, verification and validation of test methods.

- This Terrestrial Manual provides more detailed and specific guidance on test selection, optimisation, standardisation, and

- validation in chapter 1.1.6. Chapter 1.1.6 refers to chapters 2.2.1–2.2.8 Recommendations for validation of diagnostic tests

- that deal with the development and optimisation of fundamentally different assays such as antibody, antigen and nucleic

- acid detections tests, measurement uncertainty, statistical approaches to test validation, selection and use of reference

- samples and panels, validation of diagnostic tests for wildlife, and comparability experiments after changes in a validated

- test method.

- The following are key test method issues for those involved in the quality management of the laboratory.

##### 7.1.	Selection of the test method

- Valid results begin with the selection of a test method that meets the needs of the laboratory’s customers in

- addressing their specific requirements (fitness for purpose). Some issues relate directly to the laboratory, others to

- the customer.

- 7.1.1. Considerations for the selection of a test method

- i)	International acceptance;

- ii)	Scientific acceptance;

- iii)  Appropriate or current technology;

- iv)  Suitable performance characteristics (e.g. analytical and diagnostic sensitivity and specificity,

- repeatability, reproducibility, isolation rate, limits of detection, precision, trueness, and

- uncertainty);

- v)	Suitability of the test in the species and population of interest;

- vi)  Sample type (e.g. serum, tissue, milk) and its expected quality or state on arrival at the laboratory;

- vii) Test target (e.g. antibody, antigen, live pathogen, nucleic acid sequence);

- viii) Test turnaround time;

- ix)  Resources and time available for development, adaptation, evaluation;

- x)	Intended use (e.g. export, import, surveillance, screening, diagnostic, confirmatory);

- xi)  Safety factors and biocontainment requirements;

- xii) Customer expectations;

- xiii) Throughput of test Sample numbers and required throughput (automation, robot);

- xiv) Cost of test, per sample;

- xv) Availability of reference standards, reference materials and proficiency testing schemes. (See

- also chapter 2.2.6.).

##### 7.2.	Optimisation and standardisation of the test method

- Once the method has been selected, it must be set up at the laboratory. Additional optimisation is necessary, whether

- the method was developed in-house (validation) or imported from an outside source (verification). Optimisation

- establishes critical specifications and performance standards for the test process as used in a specific laboratory.

- 7.2.1. Determinants of optimisation

- i)	Critical specifications for equipment, instruments consumables, and reagents (e.g. chemicals,

- biologicals), reference standards, reference materials, and internal controls;

- ii)	Robustness – critical control points and acceptable ranges, attributes or behaviour at critical

- control points, using statistically acceptable procedures;

- iii)	Quality control activities necessary to monitor critical control points;

- iv)	The type, number, range, frequency, and arrangement of test run controls;

- v)	Criteria for non-subjective objective acceptance or rejection of a batch of test results;

- vi)	Criteria for the interpretation and reporting of test results;

- vii)  A Documented test method and reporting procedure for use by laboratory staff;

- viii) Evidence of technical competence for those who performing the test processes methods,

- authorising test results and interpreting results.

##### 7.3. Validation of the test method

- Test method validation evaluates the test for its fitness for a given use purpose by establishing test performance

- characteristics such as sensitivity, specificity, and isolation rate; and diagnostic parameters such as positive or

- negative cut-off, repeatability, reproducibility and titre of interest or significance. Validation should be done performed

- using an optimised, documented, and fixed procedure. The extent and depth of the validation process will depend on

- logistical and risk factors. It and may involve any number of activities and amount of data, with subsequent data

- analysis using appropriate statistical methods (Chapter 1.1.6.). Acknowledging diagnostic test validation science as

- a key element in the effective detection of infectious diseases, WOAH recently published a Special Issue representing

- an up-to-date compilation of the relevant standards (WOAH and non-WOAH) and guidance documents for all stages

- of diagnostic test validation and proficiency testing, including design and analysis, as well as clear, complete and

- transparent reporting of validation studies in the peer-reviewed literature (Colling & Gardner, 2021). It is important to

- note that the current version of ISO 17025:2017 specifies that personnel must be authorised to perform validation

- and related activities, which means that training in validation and verification methods, including results interpretation,

- is likely to become more important to prove competence (Colling & Gardner, 2021).

- 7.3.1. Activities that validation might include

- i)	Field or epidemiological studies, including disease outbreak investigations and testing of samples

- from infected and non-infected animals;

- ii)	Development  of  testing  algorithms  for  specific  purposes,  e.g.  surveillance,  outbreak

- investigations, etc.;

- iii)	Repeat testing in the same laboratory to establish the effect of variables such as operator,

- reagents, equipment;

- iv)	Comparison with other, preferably standard methods and with reference standards (if available);

- v)	Collaborative studies with other laboratories using the same documented method. Ideally

- organised by a reference laboratory and including testing a panel of samples of undisclosed

- composition or titre with expert evaluation of results and feedback to the participants to estimate

- reproducibility;

- vi)	Reproduction of data from an accepted standard method, or from a reputable peer-reviewed

- publication (verification);

- vii)  Experimental infection or disease outbreak studies;

- viii) Analysis of internal quality control data.

- Validation is always a balance between cost, risk, and technical possibilities. There may be cases where

- quantities such as only basic accuracy and precision can only be given determined, e.g. when the disease is

- not present in a simplified way country or region. Criteria and procedures for the correlation of test results for

- diagnosis of disease status or for regulatory action must be developed. The criteria and procedures developed

- should account for screening methods, retesting and confirmatory testing.

- Test validation is covered in chapter 1.1.6.

##### 7.4. Uncertainty of the test method

- Statistically relevant numbers of samples from infected and non-infected animals are discussed in chapter 1.1.6. test

- validation and chapter 2.2.5 statistical approaches to validation.

##### 7.4. Estimation of Measurement Uncertainty

- Measurement of Uncertainty (MU) is “a parameter associated with the result of a measurement that characterises the

- dispersion of values that could reasonably be attributed to the measure” (Eurachem, 2012). Uncertainty of

- measurement does not imply doubt about a result but rather increased confidence in its validity. It is not the equivalent

- to error, as it may be applied to all test results derived from a particular procedure.

- Laboratories must estimate the MU for each test method resulting in a quantitative measurement included in their

- scope of accreditation, and for any methods used to calibrate equipment, included in their scope of accreditation

- (ISO/IEC 17025, 2005 2017b).

- Tests can be broadly divided into two groups: quantitative (e.g. biochemical assays, enzyme-linked immunosorbent

- assays [ELISA], titrations, real-time polymerase chain reaction [PCR], pathogen enumeration, etc.); and qualitative

- (bacterial culture, parasite identification, virus isolation, endpoint PCR, immunofluorescence, etc.).

- The determination of MU is well established in quantitative measurement sciences (ANSI, 1997). It may be given as

- a numeric expression of reliability and is commonly shown as a stated range. Standard deviation (SD) and confidence

- interval (CI) are examples of the expression of MU, for example the optical density result of an ELISA expressed as

- ± n SD, where n is usually 1, 2 or 3. The confidence interval (usually 95%) gives an estimated range in which the

- result is likely to fall, calculated from a given set of test data. For quantitative measurements, example for a top-down

- or control-sample approach are provided for an antibody ELISA in chapter 2.2.4, and by the Australian government

- webpage 28. An example for a quantitative PCR (TaqMan) assay is provided by Newberry & Colling (2021).

- The ISO/IEC 17025 requirement for “quality control procedures for monitoring the validity of tests” implies that the

- laboratory must use quality control procedures that cover all major sources of uncertainty. There is no requirement

- to cover each component separately. Laboratories may establish acceptable specifications, criteria, ranges, etc., at

- critical control points for each component of the test process. The laboratory can then implement appropriate quality

- control measures at these critical points, or seek to reduce or eliminate the uncertainty effect of each component.

- 7.4.1. Potential sources of uncertainty include:

- i)	Sampling;

- ii)	Contamination;

- iii)	Sample transport and storage conditions;

- iv)	Sample processing;

- Australian Government, Department of Agriculture, Fisheries and Forestry. Worked examples of measurement uncertainty. Measurement uncertainty in veterinary diagnostic testing – DAFF (agriculture.gov.au) (accessed 15 March 2023).

- v)	Reagent quality, preparation and storage;

- vi)	Type of reference material;

- vii)  Volumetric and weight manipulations;

- viii) Environmental conditions;

- ix)	Equipment effects;

- x)	Analyst or operator bias;

- xi)	Biological variability;

- xii)  Unknown or random effects.

- Systematic errors or bias determined by validation must be corrected by changes in the method,

- adjusted for mathematically, or have the bias noted as part of the report statement.

- If an adjustment is made to a test or procedure to reduce uncertainty or correct bias then a new

- source of uncertainty is introduced (the uncertainty of the correction). This must be assessed as part

- of the MU estimate.

- The application of the principles of MU to qualitative testing is less well defined. The determination

- and expression of MU has not been standardised for veterinary (or medical, food, or environmental)

- testing laboratories, but sound guidance exists and as accreditation becomes more important,

- applications are being developed. The ISO/IEC 17025 standard recognises that some test methods

- may preclude metrologically and statistically valid calculation of uncertainty of measurement. In such

- cases the laboratory must attempt to identify and estimate all the components of uncertainty based

- on knowledge of the performance of the method and making use of previous experience, validation

- data, internal control results, etc.

- Many technical organisations and accreditation bodies (e.g. AOAC International, ISO, NATA, A2LA,

- Standards Council of Canada, UKAS, Eurachem, the Cooperation of International Traceability in

- Analytical Chemistry) teach courses or provide guidance on MU for laboratories seeking

- accreditation.

- The ISO/IEC 17025 requirement for “quality control procedures for monitoring the validity of tests”

- implies that the laboratory must use quality control procedures that cover all major sources of

- uncertainty. There is no requirement to cover each component separately. Laboratories may

- establish acceptable specifications, criteria, ranges, etc., at critical control points for each component

- of the test process. The laboratory can then implement appropriate quality control measures at these

- critical points, or seek to reduce or eliminate the uncertainty effect of each component. Measurement

- Uncertainty is covered in chapter 2.2.4.

- 7.4.1.	Components of tests with sources of uncertainty include:

- i)	Sampling;

- ii)	Contamination;

- iii)	Sample transport and storage conditions;

- iv)	Sample processing;

- v)	Reagent quality, preparation and storage;

- vi)	Type of reference material;

- vii)  Volumetric and weight manipulations;

- viii) Environmental conditions;

- ix)	Equipment effects;

- x)	Analyst or operator bias;

- xi)	Biological variability;

- xii)  Unknown or random effects.

- Systematic errors or bias determined by validation must be corrected by changes in the method,

- adjusted for mathematically, or have the bias noted as part of the report statement.

- If an adjustment is made to a test or procedure to reduce uncertainty or correct bias then a new

- source of uncertainty is introduced (the uncertainty of the correction). This must be assessed as part

- of the MU estimate.

- Additional information on the analysis of uncertainty may be found in the Eurachem Guides to

- Quantifying Uncertainty in Measurement (Eurachem, 2012) and Use of uncertainty information in

- compliance assessment Uncertainty Information in Compliance Assessment (Eurachem, 2007).

##### 7.5. Implementation and use of the test method

- Training should be a planned and structured activity with steps to ensure adequate supervision is maintained while

- analysts are being trained. Depending on the complexity of the test and the experience of the analyst, training may

- include any combination of reading and understanding the documented test method, initial demonstration,

- performance of the test under supervision and independent performance. Analysts should be able to demonstrate

- proficiency in using the test method prior to producing being authorised to produce reported results, and on an

- ongoing basis.

- The laboratory must be able to demonstrate traceability for all accredited tests and the principle should apply to all

- tests whether accredited or not. This covers all activities relating to test selection, development, optimisation,

- standardisation, validation, verification, implementation, reporting, personnel, quality control and quality assurance

- (see also Section 7.3.1. point vi). Traceability is achieved by using appropriate documented project management,

- record keeping, data management and archiving systems.

#### 8.	Strategic planning

- Laboratories should have evidence of continual improvement, which is an obligatory requirement for accredited

- laboratories. The laboratory must be knowledgeable of and stay maintain current with knowledge of the relevant quality

- and technical management standards and with methods used to demonstrate laboratory competence and establish and

- maintain technical validity. Evidence of this may be provided by include:

- i)	Attendance at conferences, organisation of in-house or external meetings on diagnostics and quality

- management;

- ii)	Participation in Membership of local and international organisations;

- iii)	Participation in writing Contribution to national and international standards (e.g. on ILAC and ISO committees);

- iv)	Maintenance of current awareness of publications, writing through review of and reviewing publications about

- diagnostic methods contribution to relevant literature;

- v)	Participation in training programmes, including visits to other laboratories;

- vi)	Conducting research;

- vii)  Participation in cooperative programmes (e.g. Inter-American Institute for Cooperation in Agriculture);

- viii) Exchange of procedures, methods, reagents, samples, personnel, and ideas;

- ix)	Planned, continual professional development and technical training;

- x)	Management reviews;

- xi)	Analysis of customer feedback;

- xii)  Root cause analysis of anomalies and implementation of corrective, preventive and improvement actions, as

- well as effectiveness reviews.

### REFERENCES

- ANSI (AMERICAN NATIONAL STANDARDS INSTITUTE) (1997). ANSI/NCS 29 Z540-2-1997 (R2012), US Guide to the Expression

- of Uncertainty in Measurement, First Edition. American National Standards Institute, www.ansi.org.

- COLLING A.& GARDNER I.A., EDS (2021). Diagnostic test validation science: a key element for effective detection and control

- of	infectious	animal	diseases.	Rev.	Sci.	Tech.	Off.	Int.	Epiz.,	40,	Available	at

- https://doc.woah.org/dyn/portal/index.xhtml?page=alo&aloId=41245

- EURACHEM (2007). Eurachem/CITAC Guide: Use of Uncertainty Information in Compliance Assessment, First Edition.

- Eurachem Secretariat. www.eurachem.org.

- EURACHEM (2012). Eurachem/CITAC 30 Guide (CG4): Quantifying Uncertainty in Analytical Measurement, Third Edition.

- Eurachem Secretariat. www.eurachem.org.

- ISO (2015). ISO 9001:2015. Quality Management Systems – Requirements. International Organization for Standardization

- (ISO), www.iso.org.

- ISO/IEC (2004a 2020). ISO/IEC 17000:2004 2020. Conformity Assessment – Vocabulary and General Principles.

- International Organization for Standardization (ISO), www.iso.org.

- ISO/IEC (2004b 2017a). ISO/IEC 17011:20042 017. Conformity Assessment – General Requirements for accreditation

- Bodies Accrediting Conformity Assessment Bodies. International Organization for Standardization (ISO), www.iso.org.

- ISO/IEC (2005 2017b). ISO/IEC 17025:2005 2017. General Requirements for the Competence of Testing and Calibration

- Laboratories. International Organization for Standardization (ISO), www.iso.org.

- ISO/IEC (2010). ISO/IEC 17043:2010. Conformity Assessment – General Requirements for Proficiency Testing.

- International Organization for Standardization (ISO), www.iso.org

- ISO/IEC (2022 2012). ISO/IEC 15189:2012 2022. Medical Laboratories – Requirements for Quality and Competence.

- International Organization for Standardization (ISO), www.iso.org

- JOHNSON P. & CABUANG L. (2021). Proficiency testing and ring trials. Rev. Sci. Tech. Off. Int. Epiz., 40, 189–203.

- NEWBERRY K. & COLLING A. (2021). Quality standards and guidelines for test validation for infectious diseases in veterinary

- laboratories. Rev. Sci. Tech. Off. Int. Epiz., 40, 227–237.

- STEVENSON M., HALPIN K. & HEUER C. (2021). Emerging and endemic zoonotic diseases: surveillance and diagnostics. Rev.

- Sci. Tech. Off. Int. Epiz., 40, 119–129.

- WAUGH C. & CLARK G. (2021). Factors affecting test reproducibility among laboratories. Rev. Sci. Tech. Off. Int. Epiz., 40,

- 131–141.

432	*

433	* *

434	NB: FIRST ADOPTED IN 1996 AS GOOD LABORATORY PRACTICE, QUALITY CONTROL AND QUALITY ASSURANCE.

435	Most recent updates adopted in 2017.

- NCSL: The National Conference of Standards Laboratories.

- CITAC: The Cooperation of International Traceability in Analytical Chemistry.