Annex 6. Item 5.1. – Chapter 2.2.4 Measurement uncertainty

2	MEETING OF THE WOAH BIOLOGICAL STANDARDS COMMISSION

3	Paris, 4–8 September 2023

4

5	CHAP T E R  2 . 2 . 4 .

# 6	M E A S U RE M EN T U N CE RT A I N T Y

## 7	INTRODUCTION

- The WOAH Validation Recommendations provide detailed information and examples in support of the

- WOAH Validation Standard that is published as Chapter 1.1.6 Principles and methods of Validation of

- diagnostic assays for infectious diseases of terrestrial animals this Terrestrial Manual, or Chapter 1.1.2 of

- the Aquatic Manual. The Term “WOAH Validation Standard” in this chapter should be taken as referring to

- those chapters.

- Estimation of measurement uncertainty (MU), sometimes termed measurement imprecision, is a

- requirement for testing laboratories based on international quality standards such as ISO/IEC 17025-2005,

- 2017 General requirements for the competence of testing and calibration laboratories (ISO/IEC 17025). The

- measurement process for detection of an analyte in a diagnostic sample is not entirely reproducible and

- hence there is no exact value that can be associated with the measured analyte. Therefore, the result is

- most accurately expressed as an estimate together with an associated level of imprecision level. This

- imprecision is the measurement uncertainty (MU). MU is limited to the measurement process of quantitative

- tests. The approach described here is known as “top-down” or “control sample” because it uses a weak

- positive control sample and expresses the MU result at the cut-off, where it most matters. It is not a question

- of whether the measurement is appropriate and fit for whatever use to which it may be applied. It is not an

- alternative to test validation but is rightly considered a component of that process (see the WOAH Validation

- Standard, chapter 1.1.6 Section B.1.1 Repeatability).

## 25	A. THE NECESSITY OF DETERMINING MU

- To assure compliance with ISO/IEC 17025-2005 2017 requirements, national accreditation bodies for diagnostic testing

- laboratories require laboratories to calculate MU estimates for accredited test methods that produce quantitative results,

- e.g. optical densities (OD), percentage of positivity or inhibition (PP, PI), titres, cycle threshold (CT) values, etc. This

- includes tests where numeric results are calculated and then are expressed as a positive or negative result at a cut-off

- value. For the purpose of estimating MU in serology and reverse transcriptase polymerase chain reaction (RT-PCR),

- suitable statistical measures are mean target values ± 2 standard deviations (SD), which is approximately equal to a 95%

- confidence interval (CI), relative standard deviation (RSD = SD / mean of replicates) and coefficient of variation (CV = RSD

- × 100%). Examples provided below assume normal distribution of data. The concept of MU does not apply to strictly binary

- (qualitative) results (positive or negative).

#### 1.	Samples for use in determining MU

- Repeatability is the level of agreement between results of replicates of a sample both within and between runs of the same

- test method in a given laboratory. During assay development, repeatability is estimated by evaluating variation in results

- of independent replicates from a minimum of three (preferably five) samples representing analyte activity within the

- operating range of the assay (see the WOAH Validation Standard, Chapter 1.1.6 Validation of diagnostic assays for

- infectious diseases of terrestrial animals, Sections A.2.5 Robustness and B.1.1 Repeatability, and Chapter 2.2.6 Selection

- and use of reference samples and panels, Section 3.1 A.4.2). Typically, the variation in replicate results is expressed as

- RSD or CV. The significant feature is that repeatability studies can be used to define the expected precision of the assay

- in the detection of a range of analyte concentrations.

- The use of internal quality or process controls over a range of expected results has become part of daily quality control

- and quality assurance operations of accredited facilities (see the WOAH Validation Standard, chapter 1.1.6, Sections A.2.6

- Calibration of the assay to standard reagents and B.5.1 Monitoring the assay, and Chapter 2.2.6, Section 1.4 C.1). These

- results provide a continuous monitor relative to different aspects of repeatability, e.g. intra- and inter-assay variation, intra-

- and inter-operator variation and intra- and inter-batch variation, which, when subjected to statistical analysis, provide an

- expression of the level of robustness (precision) of a test procedure. The monitoring of assay quality control parameters

- for repeatability provides evidence that the assay is or is not performing as expected. For control samples to provide valid

- inferences about assay precision, they should be treated in exactly the same way as test samples in each run of the assay,

- e.g. including sample preparation such as extraction steps or dilution of serum samples for an antibody enzyme-linked

- immunosorbent assay (ELISA).

- The variation of the results for control samples can also be used as an estimate of those combined sources of uncertainty

- and is called the “top-down” approach. This approach recognises that the components of precision will be manifest in the

- ultimate measurement. So monitoring the precision of the measurement over time will effectively show the combined effects

- of the imprecision associated with component steps.

- The imprecision or uncertainty of the measurement process associated with a test result becomes increasingly more

- important the closer the test value is to the diagnostic cut-off value. This is because an interpretation is made relative to

- the assay threshold regarding the status of the test result as positive, negative, or inconclusive (as will be described in the

- following example). In this context, low weak positive samples, like those used in repeatability studies or as the low weak

- positive control, are most appropriate for estimation of MU. The rationale being that MU, which is a function of assay

- precision, is most critical at decision-making points (i.e. thresholds or cut-offs), which are usually near the lower limit of

- detection for the assay. In this chapter, the application of MU with respect to cut-off (threshold) values, whether

- recommended by test-kit manufacturers or determined in the diagnostic laboratory, is described.

- MU, using the top-down approach, ideally requires long-term accumulated data from a weak positive control sample after

- multiple test runs over time, with multiple operators and variable conditions. The examples given below are based on 10

- data points but higher numbers will increase robustness.

#### 2.	Example of MU calculations in ELISA serology

- For most antibody detection tests, it is important to remember that the majority of tests are measurements of antibody

- activity relative to a threshold against which a dichotomous interpretation of positive or negative is applied. This is important

- because it helps to decide where application of MU is appropriate. In serology, uncertainty is frequently most relevant at

- the threshold between positive and negative determinations. Results falling into this zone are also described as

- intermediate, inconclusive, suspicious or equivocal (see the WOAH Validation Standard, chapter 1.1.6, Section B.2.4

- Selection of a cut-off (threshold) value for classification of test results).

- A limited data set from a competitive ELISA for antibody to avian influenza virus is used as an example of a “top-down”

- approach for serology. A low weak positive control sample was used to calculate MU at the cut-off level 32.

##### 2.1.	Method of expression of MU

- As the uncertainty is to be estimated at the threshold, which is not necessarily the reaction level of the low weak

- positive control serum, the relative standard deviation (RSD), or coefficient of variation (CV), if expressed as a

- percentage, provides a convenient transformation:

_

- The Australian Government, Department of Agriculture, Fisheries and Forestry, has compiled worked examples for a number of diagnostic tests Available online at: https://www.agriculture.gov.au/agriculture-land/animal/health/laboratories/tests/measurement- uncertainty (accessed 22 June 2023)

| RSD | (X) | = | SD | (X) | / (X) |
| --- | --- | --- | --- | --- | --- |

- To simplify assessment, the transformed result is regarded as the assay output result, which is the averaged across

- the number of replicates (�X). In the case of this example, a competitive ELISA, results are “normalised” (as defined

- in the WOAH Validation Standard, chapter 1.1.6, Section A.2.7 ‘Normalising’ test results to a working standard) to a

- working standard by forming a ratio of all optical density (OD) values to the OD result of a non-reactive (negative)

- control (ODN). This ratio is subtracted from 1 to set the level of antibody activity on a positive correlation scale; the

- greater the level, the greater the calculated value. This adjusted value is expressed as a per cent and referred to as

- the percentage inhibition or PI value. So for the low weak positive control serum (ODL), the transformation to obtain

- the per cent inhibition values for the low weak positive control (PIL) is:

90	PIL = 100 × [1– {ODL/ ODN}]

91	The relative standard deviation becomes:

92

##### 2.2.	Example

RSD (PIL) = SD (PIL)/ (PIL)

- A limited data set for the AI competitive ELISA example is shown below. In the experiment, the operator tested the

- low weak positive control serum ten times in the same run. Ideally in the application of this “top down” method, a

- larger data set would be used, which would enable accounting for effects on precision resulting from changes in

- operator and assay components (other than only the control serum).

- Table 1. Top-down or control sample approach for an influenza antibody C-ELISA

| Test | Pl (%) |
| --- | --- |
| 1 | 56 |
| 2 | 56 |
| 3 | 61 |
| 4 | 64 |
| 5 | 51 |
| 6 | 49 |
| 7 | 59 |
| 8 | 70 |
| 9 | 55 |
| 10 | 42 |

- Mean PI = 56.3; Std Dev (SD) = 7.9; Assays (n) = 10

##### 2.3.	Calculating uncertainty

- From the limited data set,

- RSD (PIL) = SD/Mean 7.9/56.3 = 0.14 (or as coefficient of variation = 14%)

- Expanded uncertainty (U) is the statistic defining the interval within which the value of the measure and is believed

- to lie within a specified level of confidence, usually 95%. Expanding the uncertainty is done by multiplying the RSD

- (PIL) by a factor of 2; this allows the calculation of an approximate 95% confidence interval around the threshold value

- (in this case at PI = 50%), assuming normally distributed data.

- U (95%CI) = 2 × RSD = 0.28

- This estimate can then be applied at the threshold level

- 95% CI = 50 ± (50 × 0.28) = 50 ± 14%

##### 2.4.	Interpretation

- Any positive result (PI > 50%) that is less than 64% is not positive with 95% confidence. Similarly, a negative result

- (PI < 50%) that is higher or equal to a PI of 36 is not negative at the 95% confidence level. This zone of lower

- confidence may correlate with the “grey zone” or “inconclusive/suspect zone” for interpretation that should be

- established for all tests (Greiner et al., 1995).

- 3.	Example of MU calculation in molecular tests

##### 3.1.	Example

- For real-time PCRs, replicates of positive controls with their respective cycle threshold (CT) values can be used to

- estimate MU using the top-down approach (Newberry & Colling, 2021). The method of expression follows the same

- formula as for the ELISA example above. This example uses data from replicate runs of a weak positive control

- sample (10 runs) of an equine influenza hydrolysis probe assay.

- Table 2. Top-down or control sample approach for an equine influenza TaqMan A assay

| Test | Ct value |
| --- | --- |
| 1 | 33.60 |
| 2 | 33.20 |
| 3 | 33.96 |
| 4 | 33.18 |
| 5 | 33.96 |
| 6 | 32.72 |
| 7 | 33.57 |
| 8 | 33.45 |
| 9 | 32.80 |
| 10 | 33.20 |

- Mean = 33.36; Std Dev (SD) = 0.43; Assay n=10

##### 3.2.	Calculating uncertainty

- From the limited data set,

- RSD (PIL) = SD/Mean 0.43/33.36 = 0.0128 (or as coefficient of variation = 1.28%)

- Expanded uncertainty (U) is the statistic defining the interval within which the value of the measure and is believed

- to lie within a specified level of confidence, usually 95%. Expanding the uncertainty is done by multiplying the RSD

- (PIL) by a factor of 2; this allows the calculation of an approximate 95% confidence interval around the threshold value

- (in this case at Ct value = 37), assuming normally distributed data.

- U (95%CI) = 2 × RSD = 0.0255

- This estimate can then be applied at the threshold level

- 95% CI = 37± (37 × 0.0255) = 37 ± 0.94

- The mean cycle threshold (Ct) value after 10 runs is 33.36 and the standard deviation is 0.43. The relative standard

- deviation is 0.0128. The expanded uncertainty (95% CI) is 2 × the relative standard deviation = 0.0255. Measurement

- of uncertainty (MU) is most relevant at the cut-off (Ct = 37) and can be applied by multiplication (37 × 0.0255 = 0.94).

- Subtraction from the threshold (37-0.94) provides the lower 95% confidence limit (Ct = 36.06) and addition (37+0.94)

- the upper 95% confidence limit (Ct = 37.94).

##### 3.3.	Interpretation of the results

- Any positive result (Ct < 37) that is higher than 36 Ct is not positive with 95% confidence. Similarly, any negative

- result (Ct > 37) that is less than 38 Ct is not negative with 95% confidence.

## 141	B. OTHER APPLICATIONS

- The top-down approach should be broadly applicable forto a range of diagnostic tests including molecular tests. For the

- calculation of tests using a typical two-fold dilution series for the positive control such as virus neutralisation, complement

- fixation and haemagglutination inhibition tests geometric mean titre (i.e. mean and SD of log base 2 titre values) of the

- positive control serum should be calculated. Relative standard deviations based on these log scale values may then be

- applied at the threshold (log) titre, and finally transformed to represent the uncertainty at the threshold. However, in all

- cases, the approach assumes that the variance about the positive control used to estimate the RSD is proportionally similar

- at the point of application of the MU, for example at the threshold. If the RSD varies significantly over the measurement

- scale, the positive control serum used to estimate the MU at the threshold should be selected for an activity level close to

- that threshold. The Australian Government, Department of Agriculture, Fisheries and Water ResourcesForestry, has

- compiled worked examples for a number of diagnostic tests (see footnote 1). (DAFF, 2010), which are available online at:

- http://www.agriculture.gov.au/animal/health/laboratories/tests/worked-example-measurement

- For quantitative real-time PCRs (qPCR) replicates of positive controls with their respective cycle threshold (CT) values can

- be used to estimate MU using the top-down approach.

- Other approaches and variations have been described, i.e. for serological tests (Dimech et al., 2006; Goris et al., 2009;

- Toussaint et al., 2007). Additional work and policy Central documents are available from the National Pathology

- Accreditation Advisory Group and Life Science. The central document to MU isare the Guide to the expression of

- uncertainty in measurement (GUM), ISO/IEC Guide, (1995) and Eurachem/CITAC Guide, 2012 CG 4: Quantifying

- uncertainty in analytical measurement.

#### Scope and limitations of the top-down approach

- Methods for quantifying uncertainty (addressing MU) for tests vary. When estimating MU for quantitative, biologically based

- diagnostic tests, where variations in the substrate or matrix have large and unpredictable effects, a top-down approach is

- recommended (Dimech et al., 2006; Eurachem 2012; Goris et al., 2009; ISO/IEC Guide 98-3:2008; Newberry & Colling,

- 2021; Standards Council of Canada, 2021; and footnote 1). The advantage of this method is that quality control data are

- generated during normal test runs and can be used to estimate the precision of the assay and express it at the cut-off. The

- application at the cut-off depends on the performance of the test at different analyte concentrations, e.g. variation is likely

- to increase at higher diluted samples. The top-down approach does not identify individual contributors to measurement

- uncertainty but rather provides an overall estimate. Measurement uncertainty does not replace test validation; however,

- the validation process includes assessments of repeatability through quality control samples which facilitate calculation of

- MU.

### REFERENCES

- AMERICAN ASSOCIATION FOR LABORATORY ACCREDITATION (A2LA). Policy on estimating measurement uncertainty for life

- science testing labs https://portal.a2la.org/policies/A2LA_P103b.pdf (accessed 22 November 2018).

- Australian Government, Department of Agriculture and WATER RESOURCES.. Worked examples of measurement uncertainty

- http://www.agriculture.gov.au/animal/health/laboratories/tests/worked-example-measurement (accessed 22 November

- 2018).

- AUSTRALIAN GOVERNMENT DEPARTMENT OF HEALTH AND AGEING (2007). Requirements for the estimation of measurement

- uncertainty,	National	Pathology	Accreditation	Advisory	Group.

- http://www.health.gov.au/internet/main/publishing.nsf/Content/B1074B732F32282DCA257BF0001FA218/$File/dhaeou.p

- df (accessed 22 November 2018).

- DIMECH W., FRANCIS B., KOX J. & ROBERTS G. (2006). Calculating uncertainty of measurement for serology assays by use

- of precision and bias. Clin. Chem., 52, No. 3, 526–529.

- ELLISON S.L.R. & WILLIAMS A. (EDS) (2012). Eurachem/CITAC Guide CG 4: Quantifying uncertainty in analytical

- measurement,	3rd	Edition.	Eurachem,	Teddington,	United	Kingdom,	Section	3.	Available	at:

- www.eurachem.org/images/stories/Guides/pdf/QUAM2012_P1.pdf (Accessed 10 June 2023).

- GORIS N., VANDENBUSSCHE F., HERR, C., VILLERS, J., VAN DER STEDE, Y. & DE CLERCQ K. (2009). Validation of two real-time

- PCR methods for foot-and-mouth disease diagnosis: RNA-extraction, matrix effects, uncertainty of measurement and

- precision. J. Virol. Methods, 160, 157–162.

- GREINER M., SOHR D. & GOEBEL P.A. (1995). Modified ROC analysis for the selection of cut-off values and the definition of

- intermediate results of serodiagnostic tests, J. Immunol. Methods, 185, 123–132.

- ISO/IEC (1995). Guide to the expression of uncertainty in measurement (GUM), ISO/IEC Guide 98:1995. International

- Organization for Standardization (ISO), www.iso.org.

- ISO/IEC (2005). ISO/IEC 17025:20052017. General Requirements for the Competence of Testing and Calibration

- Laboratories. International Organization for Standardization (ISO), www.iso.org.

- ISO/IEC GUIDE 98-3:2008 (E). Uncertainty of measurement – Part 3 Guide to the expression of uncertainty in measurement

- (GUM:1995), International Organization for Standardization (ISO), www.iso.org (accessed 10 June 2023).

- SCC- STANDARDS COUNCIL OF CANADA. Requirements and Guidance for method Verification and Validation in Testing

- Laboratories (RG-MVVT), 2021-05-24. http://www.scc.ca/ (accessed 10 June 2023) .

- Toussaint J.F., Assam P., Caij B., Dekeyser F., Knapen K., Imberechts H., Goris N., Molenberghs G., Mintiens K., & De

- CLERCQ K. (2007). Uncertainty of measurement for competitive and indirect ELISAs. Rev. sci. tech. Off. int. Epiz., 26, 649–

- 656.

- NEWBERRY K. & COLLING A (2021). Quality standards and guidelines for test validation for infectious diseases in veterinary

- laboratories.	Rev.	Sci.	Tech.	Off.	Int.	Epiz.,	40,	227–237.	Available	at

- https://doc.woah.org/dyn/portal/index.xhtml?page=alo&aloId=41245 (accessed 22 June 2023)

205	*

206	* *

- NB: There is a WOAH Collaborating Centre for

- Diagnostic Test Validation Science in the Asia-Pacific Region (please consult the WOAH Web site:

- https://www.woah.org/en/what-we-offer/expertise-network/collaborating-centres/#ui-id-3).

- Please contact the WOAH Collaborating Centre for any further information on validation.

- NB: First adopted in 2014.