  - Annex 7. Item 5.1. – Chapter 2.2.6. Selection and use of reference samples and panels

  - MEETING OF THE WOAH BIOLOGICAL STANDARDS COMMISSION

  - Paris, 4–8 September 2023

4

USA COMMENTS IN RED FONT

CHAP T E R  2 . 2 . 6 .

# 6	SEL ECT I ON A N D U SE OF

7	R EF ER EN CE SA M PLES AN D PA N EL S

## 8	INTRODUCTION

- The WOAH Validation Recommendations provide detailed information and examples in support of the

- WOAH Validation Standard that is published as Chapter 1.1.6 Principles and methods of Validation of

- diagnostic assays for infectious diseases of terrestrial animals this Terrestrial Manual, or Chapter 1.1.2 of

- the Aquatic Manual. The Term “WOAH Validation Standard” in this chapter should be taken as referring to

- those chapters.

- Reference samples and panels are essential from the initial proof of concept in the development laboratory

- through to the maintenance and monitoring of assay performance in the diagnostic laboratory and all of the

- stages in between. The critical importance of reference samples and panels cannot be over-emphasised.

- The wrong choice of reference materials can lead to bias and flawed conclusions right from development

- through to validation and use. Therefore, care must be exercised in selecting reference samples and

- designing panels.

- Fig. 1. Reference samples and panels grouped based on similar characteristics and composition. The

- topics and alphanumeric subheadings (e.g. Proof of concept, A.2.1) refer to the relevant section in the

- WOAH Validation Standard, Chapter 1.1.6 Validation of diagnostic assays for infectious diseases of

- terrestrial animals.

| Group A |  | Group B |  | Group D |
| --- | --- | --- | --- | --- |
| Proof of concept, A.2.1. |  | Asp, B.1.2. |  | Standard method comparison, B.2.6. |
| Operating range, A.2.2 3. |  | Analytical accuracy, ancillary tests B.1.4. |  | Provisional recognition, B.2.6 7. |
| ASe, B.1.3. |  | Reference samples and panels |  | Biological modifications, B.5.2.2. |
| Optimisation, A.2. 3 2. |  | Group C |  | Group E |
| Robustness, A.2.5. Preliminary repeatability, A.2.8. |  | Repeatability B.1.1. |  | DSp and DSe Gold standard, B.2.1. |
| Calibration and process control, A.2.6. |  | Preliminary reproducibility, B.2.6 7. |  | Group F |
| Process control, A.2.6. |  | Reproducibility, B.3. |  | DSp and DSe no gold standard B.2.2. |
| ASe, B.1.3. |  | Proficiency testing, B.5.1. |  |  |
| Technical modifications, B.5.2.1. |  |  |  |  |
| Reagent replacement, B.5.2.3. |  |  |  |  |

- ASp = Analytical specificity; ASe = Analytical sensitivity; DSp = diagnostic specificity; DSe = diagnostic sensitivity

- As can be seen in Figure 1, Reference samples and/or panels are mentioned throughout the WOAH

- Validation Standard, chapter 1.1.6. As defined in the glossary of the OIE Quality Standard and Guidelines

- for Veterinary Laboratories: Infectious Diseases, ‘Reference materials are “substances whose properties are

- sufficiently homogenous and well established to be used for the calibration of an apparatus, the assessment

- of a measurement method, or for assigning values to materials” 33. In the context of test method validation,

- reference materials or samples contain the analyte of interest in varying concentrations or activities

- reactivities and are used in developing and evaluating the candidate assay’s analytical and diagnostic

- performance characteristics. In our case, . Analyte means the specific component of a test sample that is

- detected or measured by the test method,  antibody, antigen or nucleic acid. These Reference samples

- may be sera, fluids, tissues, excreta, feed and or environmental samples that contain the analyte of interest

- and are usually harvested from infected animals and their environments. However, in some cases, they may

- be prepared in the laboratory from an original starting material ( a dilution of a high positive serum in

- negative serum) or perhaps created by spiking the chosen matrix with a derived analyte ( a bacterial or

- viral culture, a recombinant/expressed protein, or a genomic construct). Whether natural or prepared, they

- are used in experiments throughout the development process, carry over into the validation pathway and

- can be used to monitor performance throughout the lifespan of the assay.

- In Figure 1, reference samples and panels are grouped based on similar characteristics and composition

- and these groupings will be the basis for the following descriptions. As a cross-reference, the appropriate

- Section of the OIE Validation Standard is indicated under each particular application of the reference sample

- or panel.

- Reference samples may be used for multiple purposes from the initial stages of development and

- optimisation, through Stage 1 and into continual monitoring and maintenance of the assay. Wherever

- possible, large quantities of these reference samples should be collected or prepared and preserved for

- long-term use. Switching reference samples during the validation process introduces an intractable variable

- that can severely undermine interpretation of experimental data and therefore, the integrity of the

- development and validation process. For assays that may target multiple species, the samples should be

- representative of the primary species of interest. It is critical that these samples reflect both the target analyte

- and the matrix in which it is found in the population for which the assay is intended. The reference materials

- should appropriately represent the range of analyte concentration to be detected by the assay.

- It is important to emphasise that, no matter Whether reference samples are selected from natural sources

- or prepared in the laboratory, all selection criteria or and preparation procedures, as well as testing

- requirements, need to be fully described and put into document control. Not only is this good quality

- management practice, but it will provide both an enhanced level of continuity and confidence throughout the

- lifespan of the assay. Summaries of the data to be collected and documented for reference material can be

- found in Figure 2. For more detail on best practice and quality standards for the documentation of

- provenance of reference material refer to Watson et al. (2021).

- https://www.techlab.fr/Commun/UK_Def_MRC.asp

- Figure. 2. Documentation of reference material should be thorough to ensure i) transparency of intended

- purpose during assay development; ii) the correct sample types are used in all stages of assay

- development and validation; iii) accurate replacement of depleted reagents; and iv) appropriate choice of

- reference material during assay modification and re-validation. Minimum descriptive metadata are listed for

- pathogen, animal host, tissue type and phase of infection.

| Pathogen data | Animal host and sample type data | Phase of Infection data |
| --- | --- | --- |
| - Strain/isolate - Serotype - Genotype - Lineage - Tests used for characterisation | - Natural infection - Experimental infection and protocol used - Species - Breed - Age - Sex - Reproductive status - Vaccination history - Herd history | - Clinical signs - Antibody profiles - Pathogen loading and shedding - Tests used to determine status of disease/infection (case definition) |
|  | - Tissue type/s (matrix) used - For spiked samples – detail source of analyte and diluent (matrix) used - Details relating to pooling of samples |  |

66	A. GROUP A

- The question of pooling of samples to create a reference sample is often asked. If reference material is harvested from a

- single animal, it is important to ascertain  it is representative of a typical course and stage of infection within

- the context of the population to be tested. If not, this could lead to bias and flawed conclusions related to validation. Pooling

- is a good alternative but it is imperative to pool from animals that are in a similar phase of infection. This is particularly

- important for antibody detection systems. Pooling also addresses the issue of the larger quantities of reference material to

- be stored for long term use, especially when dealing with smaller host species. Before pooling any samples, it is preferable

- that they be independently tested to demonstrate that they are similar with respect to analyte concentration and/or

- reactivity. There should be an assessment following pooling to ensure that unforeseen interference is not introduced by

- the pooling of multiple samples, for example differing blood types or antibody composition within the independent samples

- could cross-react within the pool, thus causing the pooled sample to behave differently in the specified assay than the

- individual samples when tested independently.

- It is often difficult to obtain individual samples that truly represent analyte concentrations or reactivities across the spectrum

- of the expected range. Given the dynamics of many infections or responses to pathogens, intermediate ranges are often

- very transient. In the case of antibody responses, early infection phases in individual animals often result in highly variable

- and heterogeneous populations of antibody isotypes and avidities. In general, these do not make good reference samples

- for assessing the analytical characteristics of an assay. They are nonetheless important for different types of reference

- panels as will be discussed later. For most applications in Group A, it is acceptable to use prepared samples that are

- spiked with known concentrations of analyte or a dilution series of a high positive in negative matrix to create a range of

- concentrations.

- Whether natural or prepared, reference samples should represent the anticipated range of analyte concentrations, from

- low weak to high strong positive, which would be expected during a typical course of infection. A negative reference sample

- should be included as a background monitor. If a negative (matrix) is used as diluent for preparation of a positive reference

- sample (e.g. a negative serum used to dilute a high positive serum or tissue spiked with a construct), that negative should

- definitely be included as the negative reference sample.

- As mentioned above, all reference samples should be well characterised. This includes documentation on both the

- pathogen and donor host. For pathogens, this may include details related to strain, serotype, genotype, lineage, etc. The

- source of the host material should be well described with respect to species, breed, age, sex, reproductive status,

- vaccination history, herd history, etc. Wherever possible, the phase of infection should be noted. This could include details

- related to clinical signs, antibody profiles, pathogen load or shedding, etc. Equally important, tests that are used to

- determine disease/infection status need to be well documented (see Section E of this chapter for further explanation). In

- some cases, experimental infection/exposure may be the only viable option for the production of reference material. In this

- case, all of the above considerations plus the experimental protocol should be detailed.

- Above all else, natural or prepared, reference materials must be unequivocal with respect to their status as representing

- either a true positive or a true negative sample. This may require that the status be confirmed using another test or battery

- of tests. For example, many antibody reference sera are characterised using multiple serological tests. This provides not

- only confidence but additional documented characteristics that may be required when attempting to replace or duplicate

- this reference material in the future.

- Recommendations regarding stability and storage of reference materials are available: https://www.woah.org/en/what-we-

- offer/veterinary-products/#ui-id-4

#### 1.	Proof of concept (WOAH Validation Standard, Chapter 1.1.6, Section A.2.1)

- The WOAH Validation Standard, Chapter 1.1.6states that test methods and related procedures must be appropriate for

- specific diagnostic applications in order for the test results to be of relevance. In other words, the assay must be ‘fit for

- purpose’. Many assays are developed with good intentions but without a specific application in mind. At the very outset, it

- is critical that the diagnostic purpose(s) should be defined with respect to the population(s) to be tested. The most common

- purposes are listed in broad terms in Section A of the WOAH Validation Standard, chapter 1.1.6. As such, they are inclusive

- of more narrow and specific applications. However, these specific purpose(s) need to be clearly defined from the outset

- and are critically important in the context of a fully validated assay. As will be seen in the following descriptions, clearly

- defining the application will have impact on both the selection of reference samples and panels and the design of analytical

- and diagnostic evaluations.

#### 2.	Operating range (WOAH Validation Standard, Chapter 1.1.6, Section A.2.2 3) and

- analytical sensitivity (WOAH Validation Standard, Chapter 1.1.6, Section B.1.3)

##### 2.1.	Analytical approaches Operating range and analytical sensitivity

- The operating range of the assay is defines the lower and upper analyte detection limits and the interval of analyte

- concentrations (amounts) over which the method provides suitable accuracy and precision. It also defines the lower

- and upper detection limits the assay. To establish this range, The operating range is established by serial dilution, to

- extinction, of replicates of a high strong positive reference sample is selected. This high positive sample, either natural

- or prepared, is serially diluted to extinction. Dilutions of the strong positive are made in a negative matrix

- representative of the typical sample matrix of samples type taken from animals in the population targeted by the

- assay. This includes antibody assays where a high replicates of a strong positive reference serum should be diluted

- in a negative reference serum to create the dilution series. Analytical sensitivity (ASe) is measured by replicates of

- the lower limit of detection (LOD) of an analyte in an assay. The same high strong positive reference sample may be

- used to determine both the operating range and the analytical LOD.

##### 2.2.	Comparative approaches to analytical sensitivity

- If the intended purpose is to detect low levels of analyte or subclinical infections, it may be difficult to obtain the

- appropriate reference materials from early stages of the infection process. In some cases, it may be useful to

- determine a comparative ASe by running a panel of samples on the candidate assay and on another independent

- assay. Ideally this panel of samples would be serially collected from either naturally or experimentally infected animals

- and should represent infected animals early after infection, on through to the development of clinical or fulminating

- disease, if possible. This would provide a relative comparison of ASe between the assays, as well as, and a temporal

- comparison of the earliest point of detection relative to the pathogenesis of the disease.

- An experiment like the one described above, provides a unique opportunity to collect reference samples representing

- a natural range of concentrations that would be useful for other validation purposes. Care must be taken to avoid use

- of such samples when inappropriate (consult Group D below). Wherever possible serial samples should be collected

- from at least five a statistically sound number of animals throughout the course of infection. In cases where sampling

- is lethal (e.g. requiring the harvest of internal organ tissues), the number of animals required would be a minimum

- depends on need and fitness of five per sampling event the experimental approach. In all cases approval from an

- ethics committee is required. For smaller host species, this the number may need to be increased in order to collect

- sufficient reference material. Given that experiments like this require a high commitment of resources, it would be

- wise to maximise the collection of not only the currently targeted reference samples but additional materials (e.g.

- multiple tissues, fluids, etc.) that may be useful as reference materials in the future.

#### 3.	Optimisation (WOAH Validation Standard, Chapter 1.1.6, Section A.2.32) and preliminary

- repeatability (WOAH Validation Standard, Chapter 1.1.6, Section A.2.68)

- Optimisation is the process by which the most important physical, chemical and biological parameters of an assay are

- evaluated and adjusted to ensure that the performance characteristics of the assay are best suited to the intended

- application. At least three reference samples representing negative, low weak and high strong positive may be chosen

- from either natural or prepared reference samples. Optimisation experiments are rather exhaustive especially when assays

- with multiple preparatory and testing steps are involved. It is very important that a sufficient quantity of each reference

- sample be available to complete all optimisation experiments. Changing reference samples during the course of

- optimisation is not recommended as this will result in the addition of an uncontrolled variable and a disruption in the

- continuity of optimisation evidence.

- Assessment of repeatability should begin during assay development and optimisation stages. Repeatability and is further

- verified during Stage 1 of assay validation (Section B.1.1 of chapter 1.1.6). The same reference samples should be used

- for both processes, again throughout to provide continuity of evidence.

#### 4.	Calibration and process controls (WOAH Validation Standard, Chapter 1.1.6, Section

161	A.2.6)

##### 4.1.	International, national or in-house analyte reference standards

- International reference standards are highly characterised, contain defined concentrations of analyte, and are usually

- prepared and held by international reference laboratories. They are the reagents to which all assays and/or other

- reference materials should be standardised. National reference standards are calibrated by comparison with an

- international standard reagent whenever possible. In the absence of an international standard, a national reference

- standard may be selected or prepared and it then becomes the standard of comparison for the candidate assay. In

- the absence of both of the above, an in-house standard should be selected or prepared by the development laboratory

- within the responsible organisation. In all cases, thorough documentation of reference material should be observed

- as summarised in Figure 2. All of the standard reagents, whether natural or prepared, must be highly characterised

- through extensive analysis, and preferably the methods for their characterisation, preparation, and storage have been

- published in peer-reviewed publications (Watson et al., 2021). These reference standards should also be both stable

- and innocuous.

- Reference standards, especially antibody, are usually provided in one of two formats. They may be provided as a

- single positive reagent of given titre with the expectation that the candidate assay will be standardised to give an

- equivalent titre. This is a straight forward analytical approach but many of these ‘single’ standards have been prepared

- from highly positive samples as a pre-dilution in a negative matrix in order to maximise the number of aliquots

- available. The drawback here is that there is no accounting for any potential matrix effect in the candidate assay as

- there is no matrix control provided. The other approach is to provide a negative and a low weak and high strong

- positive set of reference standards that are of known concentrations or reactivities and are within the operating range

- of the standard method that was used to prepare them. The negative provided in the set must be the same as the

- negative diluent used to prepare the weak and strong positive reference standard, if the positive standards were

- diluted. This compensates for any potentially hidden matrix effect. In addition, this set of three acts as a template for

- the selection and/or preparation of process controls (discussed below).

- Classically, the above standards usually have been polyclonal antibody standards and to a lesser extent, conventional

- antigen standards used for calibration of serological assays. However, today, reference standards could also be

- monoclonal antibodies or recombinant/expressed proteins or genomic constructs, if they are to be used to calibrate

- assays to a single performance standard.

##### 4.2.	Working standards or process controls

- Working standard reagent(s), commonly known as quality or process controls, are calibrated to international, national,

- or in-house standard reagents. They are selected or prepared in the local matrix which is found in the population for

- which the assay is intended. Ideally, negative and low weak and high strong positive working standards should be

- selected or prepared. Concentrations and/or reactivities should be within the normal operating range of the assay.

- Large quantities should be prepared, aliquoted and stored for routine use in each diagnostic run of the assay. The

- intent is that these controls should mimic, as closely as possible, field samples and should be handled and tested like

- routine samples. They are used to establish upper and lower control limits of assay performance and to monitor

- random and/or systematic variability using various control charting methods. Their daily performance will determine

- whether or not an assay is in control and if individual runs may be accepted. As such, these working reference

- samples are critically important from a quality management standpoint.

#### 5.	Technical modifications (WOAH Validation Standard, Chapter 1.1.6, Section B.5.2.1)

- Technical modifications to a validated assay such as changes in instrumentation, extraction protocols, and conversion of

- an assay to a semi-automated or fully automated system using robotics will typically not necessitate full revalidation of the

- assay. Rather, a methods comparison study may be done to determine if these minor modifications to the assay protocol

- will affect the test results. Consult See chapter 2.2.8 Comparability of assays after changes in a validated test method for

- description of experiments and statistical approaches to assay precision in the face of technical modifications that are

- appropriate for comparability testing (Bowden & Wang, 2021; Reising et al., 2021).

- In general, these approaches require the use of three reference samples, a negative, a weak and a low and high strong

- positive. Again these samples to represent the entire operating range of both assays. Samples may be either natural or

- prepared. The important point to re-iterate here is that the same reference samples that were used in the developmental

- stages of the assay may be used to assess modifications after the method has been put into routine diagnostic use. This

- provides a higher level of confidence assessing potential impacts because the performance characteristics of these

- reference samples have been well characterised. At the very least, if new reference samples are to be used, they should

- be selected or prepared using the same criteria or preparation procedures established for previous materials. Again as this

- enhances the continuity of evidence.

#### 6.	Reagent replacement (WOAH Validation Standard, Chapter 1.1.6, Section B.5.2.3)

- When a reagent such as a process control sample is nearing depletion, it is essential to prepare and repeatedly test a

- replacement before such a control is depleted. The prospective replacement should be included in multiple runs of the

- assay in parallel with the original control to establish their proportional relationship. It is important to change only one

- control reagent at a time to avoid the compound problem of evaluating more than one variable.

- Again, it cannot be over-emphasised that any Replacement reference reagent should be selected or prepared using the

- same criteria or preparation procedures established for previous materials. Again as this enhances the continuity of evidence

- and confidence in the assay and underlines the importance of documentation of reference material data (Figure 2).

| 223 |  | B. GROUP B |
| --- | --- | --- |
| 224 | 1. | Analytical specificity (WOAH Validation Standard, Chapter 1.1.6, Section B.1.2) |

- Analytical specificity (ASp) is the degree to which the assay distinguishes between the target analyte and other components

- that may be detected in the assay. This is a relatively broad definition that is often not well understood. ASp may be broken

- down into different elements as described below.

- The choice of reference samples that are required to assess ASp is highly dependent on the specific intended purpose or

- application that was originally envisaged defined at the development stage of the assay. Assessment of ASp is a crucial

- element in proof of concept and verification of fitness for purpose and may be broken down into three elements: selectivity,

- exclusivity and inclusivity.

- Selectivity: an important element is the extent to which a method can accurately detect and or quantify the targeted analyte

- in the milieu of nucleic acids, proteins and/or antibodies in the test matrix. This is sometimes termed ‘selectivity’. An

- example is the use of reference samples for tests that are designed to differentiate infected from vaccinated animals (DIVA

- tests).

- Reference samples need to be selected and tested from i) non-infected/non-vaccinated, ii) non-infected/vaccinated, iii)

- infected/non-vaccinated, and iv) infected/vaccinated animals. These samples may be collected under field conditions but

- it is important that an accurate history be collected, ideally with respect to the animals, but at least to the herds involved,

- including vaccination practices and disease occurrences (Figure 2). Alternatively, it may be necessary to produce this

- material in experiments like those described in Section A.2.2 of this chapter, but including a combination of experimentally

- vaccinated and challenged animals. It Application of the 3 Rs (replacement, reduction and refinement) aims to avoid or

- minimise the number of animals used in experiments. For enzyme-linked immunosorbent assays (ELISAs), it is important

- to avoid use of the vaccine as capture antigen in the assay (e.g. indirect ELISA enzyme-linked immunosorbent assay [I-

- ELISA]), because carrier proteins in the vaccine may stimulate non-specific antibody responses in vaccinated animals that

- may be detected in ELISA leading to false positives in the assay. Similarly to the comparative approach described above

- with respect to ASe, at least five animals in each group should be considered. For smaller host species, this number may

- need to be increased in order to collect sufficient reference material., leading to false positives in the assay. Depending on

- the DIVA test, a single experiment could be designed to assess aspects of both ASe and ASp.

- A second element, sometimes termed ‘exclusivity’, Exclusivity is the capacity of the assay to detect an analyte or genomic

- sequence that is unique to a targeted organism, and excludes all other other known organisms that are potentially cross-

- reactive. This is especially true in serological assays where there are many examples of antigens expressed by other

- organisms that are capable of eliciting cross-reacting antibody. An attempt should be made to obtain reference samples

- from documented cases of infections and/or organisms that may be cross-reactive. Depending on the type of assay, these

- reference materials may represent the organism itself, host-derived samples, or genomic sequences. A profile for the

- exclusivity of the assay should be established, and expanded on a continual basis as potentially cross-reactive organisms

- arise.

- Thirdly, a critical design consideration Inclusivity relates to the capacity of an assay to detect one or several strains or

- serovars of a species, several species of a genus, or a similar grouping of closely related organisms viruses, bacteria or

- antibodies. This defines the scope of detection and thus the fitness for purpose. Reference samples are required to define

- the scope of the assay The scope of the assay defines the choice of reference samples and the results will determine proof

- of fitness. If for example an assay is developed as a screening test to detect all known genotypes or serotypes of a virus,

- then reference samples from each representative type should be tested. As new lineages or serotype variants arise, they

- too should be tested as part of the test profile, which should be updated on an ongoing basis.

#### 2.	Analytical accuracy of adjunct ancillary tests (WOAH Validation Standard, Chapter 1.1.6,

- Section B.1.4)

- Some test methods or procedures are solely analytical tools and are usually applied used to further characterise an analyte

- that has been detected in a primary assay, for example assays like. Examples are the virus neutralisation tests used to

- type an isolated virus or characterise an antibody response and subtyping of haemagglutinin genes by polymerase chain

- reaction of avian influenza virus. Such adjunct ancillary tests must be validated for analytical performance characteristics,

- but and differ from to routine diagnostic tests because they do not require validation for diagnostic performance

- characteristics. The analytical accuracy of these tests is often dependant on the use of reference reagents material. These

- reagents, whether they are antibody for typing strains of organisms or reference strains of the organism, etc., should be

- thoroughly documented, as required for any other reference material (Figure 2), with respect to their source, identity and

- performance characteristics.

275	C. GROUP C

- Reference samples in Group C may be used for a number of purposes. In the initial development stages, they may be

- used in the assessment of assay repeatability and both preliminary reproducibility in Stage 1 and the more in depth

- assessment of reproducibility in Stage 3 of the Validation Pathway. However, these samples have a number of other

- potential uses once the assay is transferred to the diagnostic laboratory. They may be used as panels for training and

- qualifying of analysts, and for assessing laboratory proficiency in external ring testing programmes. Ideally, 20 or more

- individual samples should be prepared in large volumes. About a quarter (25%) should be negative samples and the

- remainder (75%) should represent a collection of positives spanning the operating range of the assay. They should be

- aliquoted into individual tubes in sufficient volumes for single use only and stored for long term use (Chapter 1.1.2

- Collection, submission and storage of diagnostic specimens). The number of aliquots of each that will be required will

- depend on how many laboratories will be using the assay on a routine diagnostic basis and how often proficiency testing

- is anticipated. Ideally, they should be prepared in an inexhaustible quantity, but this is seldom feasible. At a minimum,

- several hundred or more aliquots of each should be prepared at a time if the assay is intended for use in multiple

- laboratories. This allows assessment of laboratory proficiency by testing the same sample over many testing intervals – a

- useful means of detecting systematic error (bias) that may creep into long term use of an assay.

- These samples may be natural or prepared from either single or pooled starting material. The intent is that they should

- mimic as closely as possible a true test sample. Because mass storage is always a problem, it may be necessary to store

- these materials in bulk and prepare working aliquots from time to time. However, if storage space is available, it is

- preferable to prepare and store large numbers of aliquots at one time because bulk quantities of analyte, undergoing

- freeze–thaw cycles to prepare a few aliquots at a time, may be subject to degradation. Because this type of reference

- material is consumed at a fairly high rate, they will need to be replaced or replenished on a continual basis. As potential

- replacement material is identified during routine testing or during outbreaks, it is advisable to work with field counterparts

- to obtain bulk reference material and store it for future use. Alternatively, it may be necessary to produce this material in

- experiments like those described in Section A.2.2 of this chapter. Similar to the comparative approach described above

- with respect to ASe, at least five animals in each group should be considered. For smaller host species, this number may

- need to be increased in order to collect sufficient reference material.

- 1.	Repeatability (WOAH Validation Standard, Chapter 1.1.6, Section B.1.1) and preliminary

- reproducibility provisional assay recognition (WOAH Validation Standard, Chapter 1.1.6,

- Section B.2.6)

- Repeatability is the level of agreement between results of replicates of a sample both within and between runs of the same

- test method in a given laboratory. Repeatability is estimated by evaluating variation in results of replicates from a minimum

- of three (preferably five) samples representing analyte activity within the operating range of the assay. Consult Chapter

- 2.2.4 Measurement uncertainty for statistical approaches for measures of uncertainty for assessments of repeatability.

- Reproducibility is the ability of a test method to provide consistent results, as determined by estimates of precision, when

- applied to aliquots of the same samples tested in different laboratories. However, preliminary reproducibility estimates of

- the candidate assay should be determined during developmental stages. A small panel of three (but preferably five)

- representing negative, weak and both low and high strong positives, like those described above, would be adequate. This

- type of panel could also be used for a limited evaluation of reproducibility to enhance provisional acceptance status for the

- assay. The test method is usually assessed in one two or more laboratories with a high level of experience and proficiency

- in assays similar to the candidate assay. The panel of ‘blind’ samples is evaluated using the candidate assay in each of

- these laboratories, using the same protocol, same reagents and comparable equipment. This is a scaled-down version of

- Stage 3 of assay validation. Consult Chapter 2.2.4 for further explanation of the topic and its application.

#### 2.	Reproducibility (WOAH Validation Standard, Chapter 1.1.6, Section B.3)

- Reproducibility is an important measure of the precision of an assay when used in a cross-section of laboratories located

- in distinct or different regions or countries using the identical assay (protocol, reagents and controls). As the number of

- laboratories increases, so does the number of variables encountered with respect to laboratory environments, equipment

- differences and technical expertise. These An overview of the factors affecting testing reproducibility is provided in Waugh

- & Clark (2021). Reproducibility studies are a measure of an assay’s capacity to remain unaffected by substantial changes

- or substitutions in test conditions anticipated in multi-laboratory use (e.g. shipping conditions, technology transfer, reagents

- batches, equipment, testing platforms and/or environments). Each of At least three laboratories should test the same panel

- of ‘blind’ samples containing a minimum of 20 samples, representing negative and a range of positive samples. If selected

- negative and/or positive samples in the panel are duplicated, in the panel then it may be possible to assess both assay

- reproducibility and within-laboratory repeatability estimates may be augmented by replicate testing of these samples when

- used in the reproducibility studies.

#### 3.	Proficiency testing (WOAH Validation Standard, Chapter 1.1.6, Section B.5.1)

- A validated assay in routine use in multiple laboratories needs to be continually monitored to ensure uniform performance

- and provide overall confidence in test results. This is assessed through external quality assurance programmes. Proficiency

- testing is one measure of laboratory competence derived by means of an inter-laboratory comparison; implied is that

- participating laboratories are using the same (or similar) test methods, reagents and controls. Results are usually

- expressed qualitatively, i.e. either negative or positive, to determine pass/fail criteria. However, for single dilution assays,

- where semi-quantitative results provide are provided, additional data for assessment of analysis may assess non-random

- error among the participating laboratories. Refer to Johnson & Cabuang (2021) for an overview of proficiency testing and

- ring trials.

- Proficiency testing programmes are varied depending on the type of assay in use. For single dilution type assays, panel

- sizes also vary but a minimum of five samples, representing negative and both low and high positives, like those described

- above, would be adequate. Proficiency testing is not unlike a continuous form of reproducibility assessment. However,

- reproducibility, by definition, is a measure of the assay’s performance in multiple laboratories; whereas proficiency testing

- is an assessment of laboratory competence in the performance of an established and validated assay. Measurements of

- precision can be estimated for both the reproducibility and repeatability data if replicates of the same reference sample are

- included in this ‘blind’ panel. Consult Chapter 2.2.4 for further explanation of the topic and its application. vary but a

- minimum of five samples, representing negative weak and strong positives, would be adequate.

346	D. GROUP D

- Reference samples in Group D differ from the previous Groups in that each sample in the panel should be from a different

- individual animal. As indicated in Chapter 2.2.8 Comparability of assays after changes in a validated test method,

- experimental challenge studies often include repeated sampling of individual animals to determine the progression of

- disease, but this is a different objective than to comparing performance characteristics that would be associated with

- diagnostic sensitivity (DSe) and diagnostic specificity (DSp) of a test method. Serially drawn samples, taken on different

- days from the same animal, cannot be used as representative of individual animals in populations targeted by the assay,

- because such samples violate the rule of independence of samples required for such studies.

- Care must be taken in choosing the reference samples and the standard (independent) method used in this type of

- comparison to ensure that the analytes being detected (if different) demonstrate the same type of pathogenic profile in

- terms of time of appearance after exposure to the infectious agent, and relative abundance in the test samples chosen.

#### 1.	Standard method comparison and provisional recognition (WOAH Validation Standard,

- Chapter 1.1.6, Sections B.2.6 5 and B.2.6)

- There are situations where it is not possible or desirable to fulfil Stage 2 of the Validation Pathway because appropriate

- samples from the target population are scarce and animals are difficult to access (such as for exotic diseases). However,

- a small but select panel of highly characterised test samples representing the range of analyte concentration should be

- run in parallel in the candidate assay method and by a WOAH standard method, as published in the WOAH Manuals.

- Biobanks may be a useful resource in this context, providing well-characterised samples supported with metadata to

- enhance transparency and provenance of samples used in method comparisons (Watson et al., 2021). If the methods are

- deemed to be comparable (Chapter 2.2.8), and depending on the intended application of the assay, the choice may be

- made that further diagnostic validation is not required. For example, if the intended application is for screening of imported

- animals or animal products for exotic pathogens or confirmation of clinical signs, full validation beyond a test method

- comparison may not be feasible or warranted.

- Experience has shown that the greatest obstacle to continuing through Stage 2 of the Validation Pathway is the number of

- defined samples required to estimate diagnostic performance parameters with a high degree of certainty (WOAH Validation

- Standard, chapter 1.1.6, Section B.2). In some cases, provisional recognition by international, national or local authorities

- may be granted for an assay that has not been completely evaluated past analytical stages. The different rationales for

- provisional acceptance are well explained in the WOAH Validation Standard, chapter 1.1.6. In all cases however, sound

- evidence must exist for comparative estimates of DSp and DSe based on a small select panel of well-characterised

- samples containing the targeted analyte.

- Ideally, for both comparison with a standard method or provisional recognition, a panel of, for example, 60 samples could

- be assembled to ensure sufficient sample size for statistical analysis of the resulting data. This would include 30 ‘true’

- negatives and 30 ‘true’ positives. Wherever possible, the positives should reflect the range of analyte concentrations or

- activities expected in the target population. As mentioned above, each sample in this panel must represent an individual

- animal. Consult Chapter 2.2.5 for statistical approaches to determining methods comparability using diagnostic samples.

#### 2.	Biological modifications (WOAH Validation Standard, Chapter 1.1.6, Section B.5.2.2)

- There may be situations where changes to some of the biologicals used in the assay may be necessary and/or warranted.

- This may include changes to reagents themselves or a change to a different type of specimen which contains the same

- analyte as targeted in the original validated assay (e.g. from serum to saliva). At the very least, all of the analytical criteria

- of the validation pathway must be re-assessed before proceeding. If the analytical requisites are met, the remaining

- question relates to whether or not a full diagnostic validation is required. A similar approach to the above using a panel of

- 60 individual reference samples may be considered. However, in this case the original test method would be considered

- as the standard (independent) test and the modified method would be considered the candidate. Consult Chapter 2.2.5 for

- statistical approaches to determining methods comparability using diagnostic samples.

390	E. GROUP E

- Reference animals and reference samples in this Group E are well described in the WOAH Validation Standard, chapter

- 1.1.6, Section B.2.1). However, there are a few points that are worth re-iterating here.

#### 1.	‘Gold standard’ 34 – diagnostic specificity and diagnostic sensitivity (WOAH Validation

- Standard, Chapter 1.1.6, Section B.2.1)

- For conventional estimates of DSp, negative reference samples refer to true negative samples, from animals that have

- had no possible infection or exposure to the agent. In some situations, where the disease has never been reported in a

- country or limited to certain regions of a country, identification of true negative reference samples is usually not a problem.

- However, where the disease is endemic, samples such as these may be difficult to locate. It is often possible to obtain

- The term “Gold Standard” is limited to a perfect reference standard as described in the WOAH Validation Standard, chapter 1.1.6, Section B.2.1.2, and Chapter 2.2.5 Statistical approaches to validation, Introduction and Figure 1.

- these samples from regions within a large country or perhaps different countries where the disease in question does not

- occur or has either been eradicated or has never had the disease in question.

- Again For conventional estimates of DSe, positive reference samples refer to true positives. Care must be taken to ensure

- that the sample population is representative of the population that will be the target of the validated assay. It is generally

- problematic to find sufficient numbers of true positive reference animals, as determined by isolation of the organism. It may

- be necessary to resort to samples from animals that have been tested by a combination of methods that unequivocally

- classify animals as infected/exposed as discussed in the WOAH Validation Standard, chapter 1.1.6.

- The important point here is that All samples, irrespective of origin, must be documented as they would for any other

- reference sample so as to unequivocally to classify animals as infected or exposed, dependent on the fitness for purpose

- and proposed use of the test. As mentioned in Section A, and summarised in Figure 2, of this chapter, all reference samples

- should be well characterised. This includes documentation on both the pathogen and donor host. For pathogens, this may

- include details related and data documented to strain, serotype, genotype, lineage, etc. The source of the host material

- should be well described with respect to species, breed, age, sex, reproductive status, vaccination history, herd history,

- etc. Wherever possible, the phase of infection should be noted. This could include details related clinical signs, antibody

- profiles, pathogen load or shedding, etc. In some cases, experimental infection/exposure may be the only viable option

- ensure appropriate sample selection for the production of reference material (see the OIE Validation Standard, Section

- B.2.3). In this case, all of the above and the experimental protocol should be detailed intended purpose.

- Particularly relevant to these reference samples, the tests that are used to determine their so called ‘true’ disease/infection

- status need to be well documented in order to assess potential errors in estimates that may be carried over into the

- estimates for the candidate assay. Indeed, when using imperfect standard assays to define reference animal or sample

- status, the DSe and DSp performance estimates of the candidate assay may be flawed and often overestimated. Consult

- Chapter 2.2.5 for statistical considerations. Situations where a perfect reference is available for either positive or negative

- animals, and one where the reference is perfect for both are described for diagnostic test validation by Heuer & Stevenson

- (2021).

| 423 |  | F. GROUP F |
| --- | --- | --- |
| 424 | 1. | Animals of unknown status – diagnostic specificity and diagnostic sensitivity (WOAH |
| 425 |  | Validation Standard, Chapter 1.1.6, Section B.2.2) |

- Latent-class models are introduced in the WOAH Validation Standard, chapter 1.1.6. They do not rely on the assumption

- of a perfect reference (standard or independent) test but rather estimate the accuracy of the candidate test and the

- reference standard with the combined test results. Because these statistical models are complex and require critical

- assumptions, statistical assistance should be sought to help guide the analysis and describe the sampling from the target

- population(s), the characteristics of other tests included in the analysis, the appropriate choice of model and the estimation

- methods based on peer-reviewed literature. Consult Chapter 2.2.5 for statistical considerations.

- Reference populations, not individual reference samples, used in latent-class studies need to be well described. This

- includes documentation on both the pathogen and donor host. For pathogens, this may include details related to strain,

- serotype, genotype, lineage, etc., that may be circulating in the population. The source of the host material should be well

- described with respect to species, breed, age, sex, reproductive status, vaccination history, herd history, etc. as

- summarised in Figure 2. Wherever possible, the phase of infection in the populations should be noted with respect to

- morbidity or mortality events, recovery, etc.

- As a special note, if latent class models are to be used to ascribe estimates of DSe and DSp and include multiple

- laboratories in the design, it is possible to incorporate an assessment of reproducibility into the assessment. As stated

- above, statistical advice should be sought in this respect. Bayesian analysis of latent class models are complex and require adherence

- to critical assumptions. Statistical assistance should be sought to help guide the analysis and describe the sampling from

- the target population(s), the characteristics of other tests included in the analysis, the appropriate choice of model and the

- estimation methods (based on peer-reviewed literature). See chapter 2.2.5 for details and Cheung et al., 2021.

RATIONALE:  Latent class is a model; Bayesian is an analysis approach.

- FURTHER READING

- BOWDEN T.R., CROWTHER J.R. & WANG J. (2021). Review of critical factors affecting analytical characteristics of serological

- and molecular assays. Rev. Sci. Tech. Off. Int. Epiz., 40, 53–73. doi:10.20506/rst.40.1.3208.

- CHEUNG A., DUFOUR S., JONES G., KOSTOULAS P., STEVENSON M.A., SINGANALLUR N.B. & FIRESTONE S.M. (2021). Bayesian

- latent class analysis when the reference test is imperfect. Rev. Sci. Tech. Off. Int. Epiz., 40, 271–286.

- doi:10.20506/rst.40.1.3224

- JOHNSON P. & CABUANG L. (2021). Proficiency testing and ring trials. Rev. Sci. Tech. Off. Int. Epiz., 40, 189–203.

- https://doi.org/10.20506/rst.40.1.3217

- HEUER C. & STEVENSON M.A. (2021). Diagnostic test validation studies when there is a perfect reference standard. Rev. Sci.

- Tech. Off. Int. Epiz., 40, 261–270. doi:10.20506/rst.40.1.3223

- Reising M.M., Tong C., Harris B., Toohey-Kurth K.L., Crossley B., Mulrooney D., Tallmadge R.L., Schumann K.R., Lock,

- A.B. & LOIACONO C.M. (2021). A review of guidelines for evaluating a minor modification to a validated assay. Rev. Sci.

- Tech. Off. Int. Epiz., 40, 217–226. doi:10.20506/rst.40.1.3219

- WATSON J.W., CLARK G.A. & WILLIAMS D.T. (2021). The value of virtual biobanks for transparency purposes with respect to

- reagents and samples used during test development and validation. Rev. Sci. Tech. Off. Int. Epiz., 40, 253–259.

- doi:10.20506/rst.40.1.3222.

- WAUGH C. & CLARK G. (2021). Factors affecting test reproducibility among laboratories. Rev. Sci. Tech. Off. Int. Epiz., 40,

- 131–143. doi:10.20506/rst.40.1.3213	*

463	* *

464	NB: There is a WOAH Collaborating Centre for

465	Diagnostic Test Validation Science in the Asia-Pacific Region (please consult the WOAH Web site:

466	https://www.woah.org/en/what-we-offer/expertise-network/collaborating-centres/#ui-id-3).

467	Please contact the WOAH Collaborating Centre for any further information on validation

NB: First adopted in 2014.